PROJECT – 2

BY

NITHIN KUMAR LOKAYA

SUNNY SHINDE

HERAMB SAWANT

1 – NEURAL AMPLIFIER

- Amplifier Schematics

- Testbench

- Simulation Results

| Parameters | Targeted Value | Achieved Value |
| --- | --- | --- |
| Supply Voltage | 1.2V | 1.2V |
| Supply Current | 9uA | 9uA |
| Gain | 40 dB | 39.96 dB |
| Bandwidth | 7.5 kHz | 3.7 kHz |
| Input-referred Noise | 15nV/sqrt (Hz) | 2.6pV/sqrt (Hz) |
| CMRR | ≥42 dB | 47 dB |
| PSRR | ≥42 dB | 49 dB |

- Transistor Parameters

| Device Name | W/L (um) | Gm/Id | Vov |
| --- | --- | --- | --- |
| M1, M2 | 500/9 | 19 | 75mV |
| M3, M4, M5, M6 | 6.9/20 | 9 | 119mV |
| M7, M8 | 30.4/20 | 9 | 116mV |
| M9, M10 | 10/3 | 10 | 125m |
| McascN | 25/20 | 9 | 120mV |
| MsacsP | 7/6 | 10 | 120mV |

- Simulated Graphs

A – Gain and Bandwidth

B - CMRR

C - Noise

D - PSRR

2 – EEG AMPLIFIER

- Amplifier Schematics

- Simulation Results

| Parameters | Targeted Value | Achieved Value |
| --- | --- | --- |
| Supply Voltage | 1.2V | 1.2V |
| Supply Current | 128nA | 128nA |
| Gain | 40 dB | 55 dB |
| Bandwidth | 30 Hz | 31 Hz |
| Input-referred Noise | 2.4uV/sqrt (Hz) | 1.2uV/sqrt (Hz) |
| CMRR | ≥88 dB | 85 dB |
| PSRR | ≥80 dB | 83 dB |

- Amplifier Parameters

| Device Name | W/L (m) | Gm/Id | Vov |
| --- | --- | --- | --- |
| M1, M2 | 355n/1u | 22.527 | 55mV |
| M3, M4, M5, M6 | 650n/15u | 10.656 | 135mV |
| M7, M8 | 130n/11.6u | 9.36 | 150mV |
| M9, M10 | 285n/7u | 12.375 | 110.05mV |
| McascN | 145n/10.5u | 13.511 | 100mV |
| MsacsP | 140n/5u | 11.911 | 140mV |

- Simulated Graphs

A – Gain & Bandwidth

B - CMRR

C – Noise

D – PSRR